| DT1 Weight Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Place structure on weighing scale. |
| 2 | Record table below. |
| 3 | Repeat the test three times. |

Table 6

| DT1 Weight Test Results |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Requirement (kg) | Weight (kg) |  |  |  |
| Weight | 12.00 | 11.20 |  |  |  |

Table 7.

| DT1 Size Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Place the fully assembled Kinetix device on a flat and stable surface. |
| 2 | Use a measuring tape or ruler to measure the overall length, width, and height of the device. |
| 3 | Ensure measurements include all protruding components (e.g., frame edges, mounted sensors, and actuators). |
| 4 | Record the measured dimensions for length, width, and height as pictured below. |
| 5 | Compare the recorded dimensions against the design specification of 100 × 33 × 33 cm. |
| 6 | Verify that all measured dimensions are within the specified limits. |

Table 8.

| DT1 Size Test Procedure |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dimension | Length (cm) |  |  |  | Requirement (cm) |
| Length | 100 |  |  |  | ≤ 100 |
| Width | 67 |  |  |  | ≤ 33 |
| Height | 55 |  |  |  | ≤ 33 |

Table 9.

| DT3 ADL Range of Motion Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Place the fully assembled Kinetix device on a flat and stable surface. |
| 2 | Activate the device to perform each movement individually: hip adduction, hip abduction, hip flexion, knee flexion. |
| 3 | Measure the maximum angle achieved for each joint using computer-aided tracker software. |
| 4 | Record the measured angles for each movement. |
| 5 | Compare the recorded values against the ADL design specifications (D4–D9). |
| 6 | Verify that all measured joint angles fall within the specified ADL range of motion limits. |

Table 10.

| ST2 IMU Orientation Response Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Mount the IMU onto a mobile device. |
| 2 | Rotate the device into six orthogonal orientations (±X, ±Y, ±Z). |
| 3 | Record the acceleration values for each orientation. |
| 4 | Compare measured values against expected gravitational acceleration. |

Table 16.

| UT1 User Task Completion Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Brief participants on the purpose of the test and provide instructions. |
| 2 | Assign participants roles (physiotherapist and patient). |
| 3 | Instruct participants to perform the full workflow independently (device setup, strapping patient onto device, UI interaction, treatment initiation, session termination). |
| 4 | Observe user actions without intervention unless safety is compromised. |
| 5 | Record task completion time taken. |
| 6 | Fill up survey form for both roles (physiotherapist and patient) to gather feedback for improvement |
| 7 | Repeat for all participants (n = 10). |

Table 50

| UT2 Comfort and Usability Evaluation Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Allow participants to complete a treatment session using the device. |
| 2 | Provide the Device Comfort and Usability Survey to each participant. |
| 3 | Instruct participants to rate statements on ease of use, comfort, safety, and overall experience. |
| 4 | Collect qualitative feedback and comments from participants. |
| 5 | Compile and analyse survey results. |

Table 54.

| UT4 Safety and Emergency Response Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Instruct participants on STOP and HELP button functionality. |
| 2 | During the session, simulate scenarios requiring STOP and HELP activation. |
| 3 | Record whether the device halts immediately upon activation. |
| 4 | Evaluate ease of access and usability of safety controls. |

Table 59.

| UT4 Safety and Emergency Response Test Results |  |  |
| --- | --- | --- |
| Parameter | Observation |  |
| STOP Response | The device was able to stop upon activating the button |  |
| HELP Response |  |  |
| Accessibility | All testees was able to access the button on the UI with ease |  |

Table 60.

| UT5 System Integration Test Procedure |  |
| --- | --- |
| Step | Procedure |
| 1 | Conduct a full session. |
| 2 | Observe interaction between hardware, UI, and user. |
| 3 | Verify BLE connection, data display, and session logging. |
| 4 | Record any system errors or interruptions. |

Table 61.

| UT5 System Integration Test Results |  |  |
| --- | --- | --- |
| Parameter | Observation |  |
| UI–Device Connection | Stable connection |  |
| Data Display | Accurate mode, repetition and duration was displayed |  |
| Session Execution | The session was executed well |  |

Table 62.